# Response to Reviewer 2 — jimaging-4539482

Error Injection in Task-Based Image Quality Pipelines: What Regression Testing Cannot Catch, and Why Neither Internal Identities nor Closed-Form References Suffice Alone

Shuji Yamamoto

Thank you for the careful reading. Every point you raised is addressed below, in order, with the change made and where to find it. All changes are highlighted in the revised manuscript.

Thank you for the generous assessment and for nine precise requests. Almost all are places where the paper used a term without introducing it, which is a fair criticism of a paper that hopes to be read outside its own subfield. Each is answered by defining the term at first use rather than by pointing at a reference.

1. Define “detectability” and how it is evaluated; explain the 1029 number; define “phantom” and “synthetic phantom”.

All three are now defined in Section 1. Detectability is introduced as the index d′ — the separation between the observer’s responses to signal-present and signal-absent images divided by the standard deviation of that response — with the note that it is dimensionless and that d′≈2 corresponds to roughly 92% correct in a two-alternative forced choice, and that for the ideal observer it is computed in closed form rather than by simulating decisions.

The 1029 passage now explains the mechanism: where the computed NPS falls below the smallest representable number, weighting by 1/NPS multiplies those bins by an enormous factor, and the result is an arithmetic artefact assembled from bins where the “signal” is rounding error. The text also makes the point that matters: a plausible d′ is near unity, so 1029 is at least visibly wrong, and the same mechanism at milder decay produces a number that is not.

A phantom is now defined as an object of known composition imaged in place of a patient, and a synthetic phantom as one existing only as an array of numbers, generated rather than scanned, so that the answer is known analytically. The distinction carries weight throughout the paper — a synthetic phantom supplies ground truth and a physical one does not — so it is worth the sentence.

2. Define “severity” and “severity dial”. Is severity an exact parameter that can be exactly evaluated, and how? Also explain “omitted-sinc bias” and “floating-point underflow”.

Section 2.4 now says that severity is an exact parameter of the injection and not an estimate of anything: it is the interpolation weight between the correct implementation and the defective one, defined separately for each defect and computed rather than measured. A severity dial is that parameter used as a continuous control, and the text now says why that matters — a defect that is only present or absent gives one data point, whereas one that can be turned up from nothing gives a curve, and the curve is where the comparison between detection and materiality lives.

Floating-point underflow is explained where it first appears, as above. The omitted-sinc defect is defined by its formula in the same section: the bin-average and central-difference transfer functions left in the MTF estimate, MTFα=MTF⋅[sinc(fh) sinc(2fh)]α.

3. Define the edge-spread and line-spread functions and how they are formed; clarify the mean-detrending passage at lines 163–165.

Both done. The ESF is now defined at first use as the mean profile across the edge obtained by projecting every pixel onto the edge normal and binning far below the pixel pitch — which is what the edge’s tilt buys — and the LSF as its central difference between adjacent bins.

The detrending passage was genuinely confusing and has been rewritten. It now says that subtracting the mean from each region of interest sets the zero-frequency component to zero up to rounding, so the DC bin holds the residue of that subtraction rather than a measurement; that including it would make the dynamic range enormous for any field, white noise included; and that the exclusion is bookkeeping about how the estimator works, not a claim about the spectrum.

4. Define the “bridge” measure and give its mathematical expression; it is only mentioned in Table 1.

Section 2.2 now gives it in the text, as two displayed integrals — d′​2 through NEQ against the object power spectrum, and d′​2 from the prewhitening observer — with the check itself defined as the relative difference between them, and the observation that with NEQ=MTF2/NPS the two integrands are the same expression rearranged, which is why it is an identity and not an approximation.

5. Line 355, “beyond which a prewhitening observer is refused”, is unclear.

Rewritten to say what was meant: the threshold beyond which a prewhitening observer should return no value at all rather than a computed one, because 1/NPS has ceased to be numerically meaningful.

6. More detail about the Python implementation, specifically taskiq-core; and explain more clearly how AI tools were used in “code scaffolding and refactoring”.

The implementation now has a section of its own (2.7): a pure-Python package on NumPy and SciPy, containing no DICOM handling and no patient data by design — which is what allows every quantity in it to be checked against a closed form — the injection study as a single script importing the library unmodified, and the test matrix.

The second half of your comment was the more important one, and acting on it produced a correction I would not otherwise have made. The reconstruction of real projection data is not in taskiq-core at all. It is in a second package, ldct-io, which the submitted manuscript did not cite and which was not published. Please see Corrections not prompted by a comment, below.

The AI disclosure (now Section 2.8) has been made specific rather than reassuring. It states what the model drafted — initial versions of the estimator, phantom and test code, the figure and study scripts, refactoring, and manuscript text — and what it did not do: it did not choose which defects to inject, set any tolerance, or decide what the results mean.

It also states a fact that belongs in this paper more than in most: three of the six defects studied were not invented for the experiment. They were real errors in drafted code — the bin-centre position error, the radially-blurred disk, and the missing noise floor — found by the closed-form checks described here and only then turned into controlled injections. The checks were the means by which the assistance was verified, which is this paper’s argument applied to its own production.

7. Grammar, line 287: “throughout” → “throughout the interval”.

Corrected; the sentence now reads “the same estimators used throughout the study”, which is what was meant there.

8. Two suggested references.

Thank you for both. I have taken one and, following the editorial checklist’s instruction to assess suggested references critically, declined the other with a reason.

Taken. Song K-H, Shan C, Xu G, et al., Phantom evaluation of small-pixel effect in ultra-high-resolution photon-counting CT (J Appl Clin Med Phys 2026;27:e70776), now reference [21]. It is directly relevant: it reports noise texture and high-contrast resolution as the paired quantities a change of pixel size moves, which is the same coupling this paper observes under a change of reconstruction kernel. It is cited in Section 3.7 where that coupling is discussed, and used to make the point that the checks are stated on those quantities rather than on any particular detector.

Declined. Kline A, Gaonkar A, Pittman D, et al., From Redaction to Restoration: Deep Learning for Medical Image Deidentification and Reconstruction (J Digit Imaging Inform Med 2026). I read it and could not find a place where it would inform this manuscript’s argument: it concerns de-identification and restoration of medical images, whereas this paper is about implementation errors in physical and observer estimators and how they are detected. Citing it would add a reference the text does not use. I would of course include it if you see a connection I have missed.

9. Inline expressions and numerical values should be the same size as the text.

Corrected throughout. Inline mathematics is now set inline rather than as display mathematics, including the paragraph between lines 121 and 132 that you identified, and the whole text was checked for the same problem.

## Corrections not prompted by a comment

Two things were found while acting on the comments above. Both were wrong in the submitted manuscript, and I would rather report them than let them pass.

1. The Data Availability Statement was false, and the code it omitted is now published.

The submitted statement said that all the code is at taskiq-core and that every number is written by the scripts producing the figures. That is true of the synthetic arm and false of the real-scanner arm. Tables 3 and 4, Figure 5 and Sections 2.6 to 3.8 are produced by examples/acr_atlas.py in a second package, ldct-io, which supplies the DICOM-CT-PD reading, the single-slice rebinning and the filtered backprojection. That package was not cited and had never been published.

For a paper whose argument is that a result you cannot re-derive is a result you cannot check, this was not a citation slip. It has been fixed at the source rather than in the wording:

- ldct-io is now public at https://github.com/Institute-of-One/ldct-io under the MIT licence, with 60 passing tests, continuous integration, and its reconstruction validated against closed-form answers rather than against a stored output of itself;

- five example scripts had a single machine’s archive path hard-coded, which would have left a reader unable to run any of them; the location now comes from an environment variable;

- the Data Availability Statement now says which half of the work lives where, and states plainly that neither package reproduces the atlas alone.

2. The affiliation has been corrected to the form the author’s other papers carry.

The submitted manuscript omitted the postal code, and the author record entered in the submission system also reads “Institute of one, LISIT, Co., Ltd.” The correct and consistent form, as carried by this author’s papers in Journal of Imaging (12(8):392, 2026) and Tomography (12(9):125, 2026), is:

Institute of One, LISIT Co., Ltd., Tokyo 150-0044, Japan

The manuscript is corrected. I would be grateful if the author record in the submission system could be corrected to match, as I do not appear to be able to edit it at this stage.

## English language

Reviewer 1 marked the English as improvable and Reviewer 2 as not requiring improvement. I have taken the first at face value: the whole manuscript has been reread for clarity, and the passages Reviewer 2 identified as confusing — the detrending explanation, the “refused” sentence, and the definitions of severity, ESF, LSF and bridge — were the places where the writing was in fact doing the least work. Those are rewritten rather than merely polished.

Thank you both for reviews that improved the paper, and in two places corrected it.

Yours sincerely,

Shuji Yamamoto Institute of One, LISIT Co., Ltd., Tokyo 150-0044, Japan ORCID 0000-0001-9211-1071